## Test Data
Scanner and Site | FolderName | Sequence | Fully Passed? | Link to Results
Shetland Philips Ingenia 1.5T | MedACRTestingSetAndResults\ACR_HNU_Shetland\DICOM | Ax_T1_ACR_Lothian | No | Shetland Ax T1'!A1
Shetland Philips Ingenia 1.5T | MedACRTestingSetAndResults\ACR_HNU_Shetland\DICOM | Sag_T1_ACR_Lothian
Shetland Philips Ingenia 1.5T | MedACRTestingSetAndResults\ACR_HNU_Shetland\DICOM | Cor_T1_ACR_Lothian
Forth Valley SIGNA | MedACRTestingSetAndResults\ACR Blair T1 | ACR_Axial_T1 | Yes | ForthValley Ax T1'!A1
Gartnaval Philips Ingenia 1.5T | MedACRTestingSetAndResults\Blair Gartnavel | ACR_ax_T1
Raigmore GE SIGNA Artist MRI 1 | MedACRTestingSetAndResults\Raigmore ACR MRI1 Test Data | ACR AxT1
Raigmore GE SIGNA Artist MRI 1 | MedACRTestingSetAndResults\Raigmore ACR MRI1 Test Data | ACR AxT2
Raigmore GE SIGNA Artist MRI 2 | MedACRTestingSetAndResults\Raigmore ACR MRI 2 Test data | Ax T1 SE
Raigmore GE SIGNA Artist MRI 2 | MedACRTestingSetAndResults\Raigmore ACR MRI 2 Test data | Ax T2 SE
Raigmore GE SIGNA Artist MRI 2 | MedACRTestingSetAndResults\Raigmore ACR MRI 2 Test data | Sag T1 SE
Raigmore GE SIGNA Artist MRI 2 | MedACRTestingSetAndResults\Raigmore ACR MRI 2 Test data | Sag T2 SE
Raigmore GE SIGNA Artist MRI 2 | MedACRTestingSetAndResults\Raigmore ACR MRI 2 Test data | Cor T1 SE
Raigmore GE SIGNA Artist MRI 2 | MedACRTestingSetAndResults\Raigmore ACR MRI 2 Test data | Cor T2 SE

## ForthValley Ax T1
Test Case ID | ForthValley V1 |  | Test Description | Initial Overall Test
Git Commit | b0a868e
Tested By | J Tracey
Test Date | 2024-09-12
|  |  | Scanner | GE SIGNA Artist
|  |  | Scan Date | 2024-03-14
|  |  | Series Description | ACR_Axial_T1
Test | Comparable Measurement | Software Result | Manual Result | Difference | Tolerance | PASS/ FAIL | Comments
Uniformity | Integral Uniformity | 96.22% | 96.42% | 0.2074% | 1% | PASS
Slice Position | Distance Slice 1(mm) | -0.39 | -0.36 | 0.03 | 0.6 | PASS
| Distance Slice 11(mm) | -1.17 | -1.09 | 0.08 | 0.6 | PASS
Slice Width | Distance (mm) | 5.1 | 5.16 | 0.06 | 0.6 | PASS
SNR | SNR | 337.28 | 330 | 2.206% | 3% | PASS
| Normalised SNR | 4512.28 | 4412 | 2.273% | 3% | PASS
Geometric Accuracy MagNet Method | Row Top (mm) | 80.07 | 79.81 | 0.26 | 0.6 | PASS
| Row Middle (mm) | 79.67 | 80 | 0.33 | 0.6 | PASS
| Row Bottom (mm) | 79.93 | 80.2 | 0.27 | 0.6 | PASS
| Col Top (mm) | 80.11 | 80.4 | 0.29 | 0.6 | PASS
| Col Middle (mm) | 79.89 | 80 | 0.11 | 0.6 | PASS
| Col Bottom (mm) | 80.27 | 80.2 | 0.07 | 0.6 | PASS
Ghosting | Signal-to-Bkgd ratio | 0.008% | 0.01% | 2e-05 | 0.5% | PASS
Spatial Resoloution Contrast Response Manual | 1.1mm Horizontal | 34.2% | 36.64% | 2.44% | 8% | PASS
| 1.0mm Horizontal | 45% | 43.54% | 1.46% | 8% | PASS
| 0.9mm Horizontal | 0.8% | 3.01% | 2.21% | 8% | PASS
| 0.8mm Horizontal | 2.2% | 2.14% | 0.06% | 8% | PASS
| 1.1mm Vertical | 20.1% | 24.98% | 4.88% | 8% | PASS
| 1.0mm Vertical | 9.2% | 9.2% | 0% | 8% | PASS
| 0.9mm Vertical | 7% | 5.98% | 1.02% | 8% | PASS
| 0.8mm Vertical | 1.9% | 2.64% | 0.74% | 8% | PASS
Spatial Resoloution Contrast Response Automatic | 1.1mm Horizontal | 49.94% | 49.13% | 0.81% | 8% | PASS
| 1.0mm Horizontal | 40.98% | 42.13% | 1.15% | 8% | PASS | First peak not captured in range
| 0.9mm Horizontal | 7.26% | 13.01% | 5.75% | 8% | PASS | Unclear Peaks and Troughs
| 0.8mm Horizontal | 3.56% | 3.1% | 0.46% | 8% | PASS
| 1.1mm Vertical | 32.21% | 26.02% | 6.19% | 8% | PASS | Unclear Peaks and Troughs
| 1.0mm Vertical | 11.23% | 12.5% | 1.27% | 8% | PASS
| 0.9mm Vertical | 7.49% | 12.27% | 4.78% | 8% | PASS
| 0.8mm Vertical | 0.41% | 3.03% | 2.62% | 8% | PASS | Unclear Peaks and Troughs

## Shetland Ax T1
Test Case ID | Shetland Ax T1 V1 |  | Test Description | Initial Overall Test
Git Commit | b0a868e
Tested By | J Tracey
Test Date | 2024-09-12
|  |  | Scanner | Imngenia Ambition S
|  |  | Scan Date | 2024-06-04
|  |  | Series Description | Ax_T1_ACR_Lothian
Test | Comparable Measurement | Software Result | Manual Result | Difference | Tolerance | PASS/ FAIL | Comments
Uniformity | Integral Uniformity | 94.41% | 94.96% | 0.5792% | 1% | PASS
Slice Position | Distance Slice 1(mm) | 1.17 | 1.05 | 0.12 | 0.6 | PASS
| Distance Slice 11(mm) | 0.2 | 0.01 | 0.19 | 0.6 | PASS
Slice Width | Distance (mm) | 5.08 | 5.2 | 0.12 | 0.6 | PASS
SNR | SNR | 286.56 | 276.4 | 3.676% | 3% | FAIL
| Normalised SNR | 3925.42 | 3783 | 3.765% | 3% | FAIL
Geometric Accuracy MagNet Method | Row Top (mm) | 79.66 | 80.03 | 0.37 | 0.6 | PASS
| Row Middle (mm) | 79.87 | 79.72 | 0.15 | 0.6 | PASS
| Row Bottom (mm) | 79.64 | 80.01 | 0.37 | 0.6 | PASS
| Col Top (mm) | 79.72 | 79.13 | 0.59 | 0.6 | PASS
| Col Middle (mm) | 79.8 | 79.43 | 0.37 | 0.6 | PASS
| Col Bottom (mm) | 79.95 | 80.2 | 0.25 | 0.6 | PASS
Ghosting | Signal-to-Bkgd ratio | 0.023% | 0.4% | 0.00377 | 0.5% | PASS
Spatial Resoloution Contrast Response Manual | 1.1mm Horizontal |  |  | Not Run | 8% | Not Run
| 1.0mm Horizontal |  |  | Not Run | 8% | Not Run
| 0.9mm Horizontal |  |  | Not Run | 8% | Not Run
| 0.8mm Horizontal |  |  | Not Run | 8% | Not Run
| 1.1mm Vertical |  |  | Not Run | 8% | Not Run
| 1.0mm Vertical |  |  | Not Run | 8% | Not Run
| 0.9mm Vertical |  |  | Not Run | 8% | Not Run
| 0.8mm Vertical |  |  | Not Run | 8% | Not Run
Spatial Resoloution Contrast Response Automatic | 1.1mm Horizontal | 72.54% | 76.6% | 4.06% | 8% | PASS
| 1.0mm Horizontal | 32.15% | 29.75% | 2.4% | 8% | PASS
| 0.9mm Horizontal | 8.87% | 6.79% | 2.08% | 8% | PASS
| 0.8mm Horizontal | 6.06% | 9.86% | 3.8% | 8% | PASS
| 1.1mm Vertical | 74.52% | 67.6% | 6.92% | 8% | PASS
| 1.0mm Vertical | 48.26% | 44.02% | 4.24% | 8% | PASS
| 0.9mm Vertical | 11.64% | 18.31% | 6.67% | 8% | PASS
| 0.8mm Vertical | 10.95% | 8.28% | 2.67% | 8% | PASS

## Gartnavel
Test Case ID | Gart_Ax_V1 |  | Test Description | Initial Overall Test
Git Commit | b0a868e
Tested By | J Tracey
Test Date | 2024-09-12
|  |  | Scanner | Philips Ingenia 1.5T
|  |  | Scan Date | 2024-03-19
|  |  | Series Description | ACR_ac_T1
Test | Comparable Measurement | Software Result | Manual Result | Difference | Tolerance | PASS/ FAIL | Comments
Uniformity | Integral Uniformity | 97.44% | 97.02% | 0.4329% | 1% | PASS
Slice Position | Distance Slice 1(mm) | 1.56 | 1.39 | 0.17 | 0.6 | PASS
| Distance Slice 11(mm) | -0.39 | -0.58 | 0.19 | 0.6 | PASS
Slice Width | Distance (mm) | 4.94 | 5.14 | 0.2 | 0.6 | PASS
SNR | SNR | 262.09 | 251.7 | 4.128% | 3% | FAIL
| Normalised SNR | 4685.11 | 4495 | 4.229% | 3% | FAIL
Geometric Accuracy MagNet Method | Row Top (mm) | 2.93 | 79.83 | 76.9 | 0.6 | FAIL
| Row Middle (mm) | 2.93 | 79.28 | 76.35 | 0.6 | FAIL
| Row Bottom (mm) | 5.86 | 80.38 | 74.52 | 0.6 | FAIL
| Col Top (mm) | 5.86 | 80.79 | 74.93 | 0.6 | FAIL
| Col Middle (mm) | 5.86 | 79.97 | 74.11 | 0.6 | FAIL
| Col Bottom (mm) | 5.86 | 80.24 | 74.38 | 0.6 | FAIL
Ghosting | Signal-to-Bkgd ratio | 0.041% | 0.03% | 0.00011 | 0.5% | PASS
Spatial Resoloution Contrast Response Manual | 1.1mm Horizontal | 53.1% | 51.05% | 2.05% | 8% | PASS
| 1.0mm Horizontal | 16.1% | 14.59% | 1.51% | 8% | PASS
| 0.9mm Horizontal | 11.3% | 10.83% | 0.47% | 8% | PASS
| 0.8mm Horizontal | 1% | 2.25% | 1.25% | 8% | PASS
| 1.1mm Vertical | 52% | 50.85% | 1.15% | 8% | PASS
| 1.0mm Vertical | 20% | 19.83% | 0.17% | 8% | PASS
| 0.9mm Vertical | 7.4% | 7.3% | 0.1% | 8% | PASS
| 0.8mm Vertical | 4.7% | 1.19% | 3.51% | 8% | PASS
Spatial Resoloution Contrast Response Automatic | 1.1mm Horizontal | 69.21% | 71.8% | 2.59% | 8% | PASS
| 1.0mm Horizontal | 19.11% | 23.51% | 4.4% | 8% | PASS
| 0.9mm Horizontal | 4.21% | 5.05% | 0.84% | 8% | PASS
| 0.8mm Horizontal | 7.31% | 9.15% | 1.84% | 8% | PASS
| 1.1mm Vertical | 65.72% | 62.36% | 3.36% | 8% | PASS
| 1.0mm Vertical | 18.14% | 26.43% | 8.29% | 8% | FAIL
| 0.9mm Vertical | 14.2% | 6.44% | 7.76% | 8% | PASS
| 0.8mm Vertical | 11.38% | 11.51% | 0.13% | 8% | PASS

## Template
Test Case ID |  |  | Test Description
Git Commit
Tested By
Test Date
Scanner
Scan Date
Series Description
Test | Comparable Measurement | Software Result | Manual Result | Difference | Tolerance | PASS/ FAIL | Comments
Uniformity | Integral Uniformity |  |  | Not Run | 1% | Not Run
Slice Position | Distance Slice 1(mm) |  |  | Not Run | 0.6 | Not Run
| Distance Slice 11(mm) |  |  | Not Run | 0.6 | Not Run
Slice Width | Distance (mm) |  |  | Not Run | 0.6 | Not Run
SNR | SNR |  |  | Not Run | 3% | Not Run
| Normalised SNR |  |  | Not Run | 3% | Not Run
Geometric Accuracy MagNet Method | Row Top (mm) |  |  | Not Run | 0.6 | Not Run
| Row Middle (mm) |  |  | Not Run | 0.6 | Not Run
| Row Bottom (mm) |  |  | Not Run | 0.6 | Not Run
| Col Top (mm) |  |  | Not Run | 0.6 | Not Run
| Col Middle (mm) |  |  | Not Run | 0.6 | Not Run
| Col Bottom (mm) |  |  | Not Run | 0.6 | Not Run
Ghosting | Signal-to-Bkgd ratio |  |  | Not Run | 0.5% | Not Run
Spatial Resoloution Contrast Response Manual | 1.1mm Horizontal |  |  | Not Run | 8% | Not Run
| 1.0mm Horizontal |  |  | Not Run | 8% | Not Run
| 0.9mm Horizontal |  |  | Not Run | 8% | Not Run
| 0.8mm Horizontal |  |  | Not Run | 8% | Not Run
| 1.1mm Vertical |  |  | Not Run | 8% | Not Run
| 1.0mm Vertical |  |  | Not Run | 8% | Not Run
| 0.9mm Vertical |  |  | Not Run | 8% | Not Run
| 0.8mm Vertical |  |  | Not Run | 8% | Not Run
Spatial Resoloution Contrast Response Automatic | 1.1mm Horizontal |  |  | Not Run | 8% | Not Run
| 1.0mm Horizontal |  |  | Not Run | 8% | Not Run
| 0.9mm Horizontal |  |  | Not Run | 8% | Not Run
| 0.8mm Horizontal |  |  | Not Run | 8% | Not Run
| 1.1mm Vertical |  |  | Not Run | 8% | Not Run
| 1.0mm Vertical |  |  | Not Run | 8% | Not Run
| 0.9mm Vertical |  |  | Not Run | 8% | Not Run
| 0.8mm Vertical |  |  | Not Run | 8% | Not Run