## Test Data
Scanner and Site | FolderName | Sequence | Fully Passed? | Link to Results
Shetland Philips Ingenia 1.5T | MedACRTestingSetAndResults\ACR_HNU_Shetland\DICOM | Ax_T1_ACR_Lothian
Shetland Philips Ingenia 1.5T | MedACRTestingSetAndResults\ACR_HNU_Shetland\DICOM | Sag_T1_ACR_Lothian
Shetland Philips Ingenia 1.5T | MedACRTestingSetAndResults\ACR_HNU_Shetland\DICOM | Cor_T1_ACR_Lothian
Forth Valley SIGNA | MedACRTestingSetAndResults\ACR Blair T1 | ACR_Axial_T1 |  | ForthValley Ax T1'!A1
Gartnaval Philips Ingenia 1.5T | MedACRTestingSetAndResults\Blair Gartnavel | ACR_ax_T1
Raigmore GE SIGNA Artist MRI 1 | MedACRTestingSetAndResults\Raigmore ACR MRI1 Test Data | ACR AxT1
Raigmore GE SIGNA Artist MRI 1 | MedACRTestingSetAndResults\Raigmore ACR MRI1 Test Data | ACR AxT2
Raigmore GE SIGNA Artist MRI 2 | MedACRTestingSetAndResults\Raigmore ACR MRI 2 Test data | Ax T1 SE
Raigmore GE SIGNA Artist MRI 2 | MedACRTestingSetAndResults\Raigmore ACR MRI 2 Test data | Ax T2 SE
Raigmore GE SIGNA Artist MRI 2 | MedACRTestingSetAndResults\Raigmore ACR MRI 2 Test data | Sag T1 SE
Raigmore GE SIGNA Artist MRI 2 | MedACRTestingSetAndResults\Raigmore ACR MRI 2 Test data | Sag T2 SE
Raigmore GE SIGNA Artist MRI 2 | MedACRTestingSetAndResults\Raigmore ACR MRI 2 Test data | Cor T1 SE
Raigmore GE SIGNA Artist MRI 2 | MedACRTestingSetAndResults\Raigmore ACR MRI 2 Test data | Cor T2 SE

## ForthValley Ax T1
Test Case ID | Raig-01 |  | Test Description
Git Commit
Tested By | J Tracey
Test Date | 2024-09-04
|  |  | Scanner | GE SIGNA Artist
|  |  | Scan Date | 2024-03-14
|  |  | Series Description | ACR_Axial_T1
Test | Comparable Measurement | Software Result | Manual Result | Difference | Tolerance | PASS/ FAIL | Comments
Uniformity | Integral Uniformity | 96.13 | 95.45 | 0.7074% | 1% | PASS
Slice Position | Distance Slice 1(mm) | -0.39 | -0.36 |  | 0.6mm
| Distance Slice 11(mm) | -1.17 | -1.09 |  | 0.6mm
Slice Width | Distance (mm) | 5.1 | 5.16 |  | 0.6mm
SNR | SNR | 5.1 | 330 |  | 3.0% relative
| Normalised SNR | 4512.33 | 4412 |  | 3.0% relative
Geometric Accuracy MagNet Method | Horizontal CoV | 0.25 | 0.24377
| Vertical CoV | 0.24 | 0.24938
| Row Top (mm) | 80.07 | 79.81 |  | 0.6mm
| Row Middle (mm) | 79.67 | 80 |  | 0.6mm
| Row Bottom (mm) | 79.93 | 80.2 |  | 0.6mm
| Col Top (mm) | 80.11 | 80.4 |  | 0.6mm
| Col Middle (mm) | 79.89 | 80 |  | 0.6mm
| Col Bottom (mm) | 80.27 | 80.2 |  | 0.6mm
Geometric Accuracy ACR Method | Slice 1 Hor Dist (mm)
Slice 1 Vert Dist (mm)
Slice 5 Hor Dist (mm)
Slice 5 Vert Dist (mm)
Diagonal distance SW (mm)
Diagonal distance SE (mm)
Ghosting | Signal-to-Bkgd ratio | 0.008 |  |  | 0.5% absolute
Spatial Resoloution Contrast Response Automatic | 1.1mm Horizontal
1.0mm Horizontal
0.9mm Horizontal
0.8mm Horizontal
1.1mm Vertical
1.0mm Vertical
0.9mm Vertical
0.8mm Vertical
Spatial Resoloution Contrast Response Manual | 1.1mm Horizontal
1.0mm Horizontal
0.9mm Horizontal
0.8mm Horizontal
1.1mm Vertical
1.0mm Vertical
0.9mm Vertical
0.8mm Vertical