## Test Data
Scanner and Site | FolderName | Sequence | Fully Passed? | Link to Results
Shetland Philips Ingenia 1.5T | MedACRTestingSetAndResults\ACR_HNU_Shetland\DICOM | Ax_T1_ACR_Lothian | Yes | Shetland Ax T1'!A1
Forth Valley SIGNA | MedACRTestingSetAndResults\ACR Blair T1 | ACR_Axial_T1 | Yes | ForthValley Ax T1'!A1
Gartnaval Philips Ingenia 1.5T | MedACRTestingSetAndResults\Blair Gartnavel | ACR_ax_T1 | Yes | Gartnavel!A1
Raigmore GE SIGNA Artist MRI 1 | MedACRTestingSetAndResults\Raigmore ACR MRI1 Test Data | ACR AxT2 | Yes | Raig 1 ACR AxT2'!A1
Raigmore GE SIGNA Artist MRI 2 | MedACRTestingSetAndResults\Raigmore ACR MRI 2 Test data | Ax T1 SE | Yes | Raig 2 Ax T1 SE'!A1
Raigmore GE SIGNA Artist MRI 2 | MedACRTestingSetAndResults\Raigmore ACR MRI 2 Test data | Sag T1 SE | Yes | Raig 2 Sag T1 SE'!A1
Raigmore GE SIGNA Artist MRI 2 | MedACRTestingSetAndResults\Raigmore ACR MRI 2 Test data | Cor T1 SE | Yes | Raig Cor T1 SE'!A1

## ForthValley Ax T1
Test Case ID | ForthValley V1 |  | Test Description | Initial Overall Test
Git Commit | b0a868e
Tested By | J Tracey
Test Date | 2024-09-12
|  |  | Scanner | GE SIGNA Artist
|  |  | Scan Date | 2024-03-14
|  |  | Series Description | ACR_Axial_T1
Test | Comparable Measurement | Software Result | Manual Result | Difference | Tolerance | PASS/ FAIL | Comments
Uniformity | Integral Uniformity | 96.22% | 96.42% | 0.2074% | 2% | PASS
Slice Position | Distance Slice 1(mm) | -0.59 | -0.36 | 0.23 | 0.6 | PASS
| Distance Slice 11(mm) | -1.17 | -1.09 | 0.08 | 0.6 | PASS
Slice Width | Distance (mm) | 5.1 | 5.16 | 0.06 | 0.6 | PASS
SNR | SNR | 337.28 | 330 | 2.206% | 3% | PASS
Geometric Accuracy MagNet Method | Row Top (mm) | 80.07 | 79.81 | 0.26 | 0.6 | PASS
| Row Middle (mm) | 79.65 | 80 | 0.35 | 0.6 | PASS
| Row Bottom (mm) | 80.01 | 80.2 | 0.19 | 0.6 | PASS
| Col Top (mm) | 80.23 | 80.4 | 0.17 | 0.6 | PASS
| Col Middle (mm) | 79.93 | 80 | 0.07 | 0.6 | PASS
| Col Bottom (mm) | 80.28 | 80.2 | 0.08 | 0.6 | PASS
Ghosting | Signal-to-Bkgd ratio | 0.008% | 0.01% | 2e-05 | 0.5% | PASS
Spatial Resoloution Contrast Response Manual | 1.1mm Horizontal | 34.2% | 36.64% | 2.44% | 8% | PASS
| 1.0mm Horizontal | 45% | 43.54% | 1.46% | 8% | PASS
| 0.9mm Horizontal | 0.8% | 3.01% | 2.21% | 8% | PASS
| 0.8mm Horizontal | 2.2% | 2.14% | 0.06% | 8% | PASS
| 1.1mm Vertical | 20.1% | 24.98% | 4.88% | 8% | PASS
| 1.0mm Vertical | 9.2% | 9.2% | 0% | 8% | PASS
| 0.9mm Vertical | 7% | 5.98% | 1.02% | 8% | PASS
| 0.8mm Vertical | 1.9% | 2.64% | 0.74% | 8% | PASS
Spatial Resoloution Contrast Response Automatic | 1.1mm Horizontal | 49.94% | 49.13% | 0.81% | 8% | PASS
| 1.0mm Horizontal | 40.98% | 42.13% | 1.15% | 8% | PASS | First peak not captured in range
| 0.9mm Horizontal | 7.26% | 13.01% | 5.75% | 8% | PASS | Unclear Peaks and Troughs
| 0.8mm Horizontal | 3.56% | 3.1% | 0.46% | 8% | PASS
| 1.1mm Vertical | 32.21% | 26.02% | 6.19% | 8% | PASS | Unclear Peaks and Troughs
| 1.0mm Vertical | 11.23% | 12.5% | 1.27% | 8% | PASS
| 0.9mm Vertical | 7.49% | 12.27% | 4.78% | 8% | PASS
| 0.8mm Vertical | 0.41% | 3.03% | 2.62% | 8% | PASS | Unclear Peaks and Troughs

## Shetland Ax T1
Test Case ID | Shetland Ax T1 V1 |  | Test Description | Initial Overall Test
Git Commit | b0a868e
Tested By | J Tracey
Test Date | 2024-09-12
|  |  | Scanner | Imngenia Ambition S
|  |  | Scan Date | 2024-06-04
|  |  | Series Description | Ax_T1_ACR_Lothian
Test | Comparable Measurement | Software Result | Manual Result | Difference | Tolerance | PASS/ FAIL | Comments
Uniformity | Integral Uniformity | 94.41% | 94.96% | 0.5792% | 2% | PASS
Slice Position | Distance Slice 1(mm) | 1.17 | 1.05 | 0.12 | 0.6 | PASS
| Distance Slice 11(mm) | 0.2 | 0.01 | 0.19 | 0.6 | PASS
Slice Width | Distance (mm) | 5.08 | 5.2 | 0.12 | 0.6 | PASS
SNR | SNR | 286.56 | 292.05 | 1.88% | 3% | PASS
Geometric Accuracy MagNet Method | Row Top (mm) | 79.59 | 80.03 | 0.44 | 0.6 | PASS
| Row Middle (mm) | 79.94 | 79.72 | 0.22 | 0.6 | PASS
| Row Bottom (mm) | 79.63 | 80.01 | 0.38 | 0.6 | PASS
| Col Top (mm) | 79.8 | 79.81 | 0.01 | 0.6 | PASS
| Col Middle (mm) | 79.8 | 79.43 | 0.37 | 0.6 | PASS
| Col Bottom (mm) | 79.9 | 80.2 | 0.3 | 0.6 | PASS
Ghosting | Signal-to-Bkgd ratio | 0.023% | 0.4% | 0.00377 | 0.5% | PASS
Spatial Resoloution Contrast Response Manual | 1.1mm Horizontal | 54.2% | 58.4% | 4.2% | 8% | PASS
| 1.0mm Horizontal | 24.3% | 27.7% | 3.4% | 8% | PASS
| 0.9mm Horizontal | 13.7% | 14.6% | 0.9% | 8% | PASS
| 0.8mm Horizontal | 12.5% | 12.4% | 0.1% | 8% | PASS
| 1.1mm Vertical | 63.8% | 62.1% | 1.7% | 8% | PASS
| 1.0mm Vertical | 42.5% | 44.1% | 1.6% | 8% | PASS
| 0.9mm Vertical | 13.7% | 13% | 0.7% | 8% | PASS
| 0.8mm Vertical | 9.9% | 17.3% | 7.4% | 8% | PASS
Spatial Resoloution Contrast Response Automatic | 1.1mm Horizontal | 72.54% | 76.6% | 4.06% | 8% | PASS
| 1.0mm Horizontal | 32.15% | 29.75% | 2.4% | 8% | PASS
| 0.9mm Horizontal | 8.87% | 6.79% | 2.08% | 8% | PASS
| 0.8mm Horizontal | 6.06% | 9.86% | 3.8% | 8% | PASS
| 1.1mm Vertical | 74.52% | 67.6% | 6.92% | 8% | PASS
| 1.0mm Vertical | 48.26% | 44.02% | 4.24% | 8% | PASS
| 0.9mm Vertical | 11.64% | 18.31% | 6.67% | 8% | PASS
| 0.8mm Vertical | 10.95% | 8.28% | 2.67% | 8% | PASS

## Gartnavel
Test Case ID | Gart_Ax_V1 |  | Test Description | Initial Overall Test
Git Commit | b0a868e
Tested By | J Tracey
Test Date | 2024-09-12
|  |  | Scanner | Philips Ingenia 1.5T
|  |  | Scan Date | 2024-03-19
|  |  | Series Description | ACR_ac_T1
Test | Comparable Measurement | Software Result | Manual Result | Difference | Tolerance | PASS/ FAIL | Comments
Uniformity | Integral Uniformity | 97.44% | 97.02% | 0.4329% | 2% | PASS
Slice Position | Distance Slice 1(mm) | 1.56 | 1.39 | 0.17 | 0.6 | PASS
| Distance Slice 11(mm) | -0.39 | -0.58 | 0.19 | 0.6 | PASS
Slice Width | Distance (mm) | 4.94 | 5.14 | 0.2 | 0.6 | PASS
SNR | SNR | 262.09 | 266.74 | 1.743% | 3% | PASS
Geometric Accuracy MagNet Method | Row Top (mm) | 79.94 | 79.83 | 0.11 | 0.6 | PASS
| Row Middle (mm) | 79.3 | 79.28 | 0.02 | 0.6 | PASS
| Row Bottom (mm) | 79.66 | 80.15 | 0.49 | 0.6 | PASS
| Col Top (mm) | 80.11 | 80.38 | 0.27 | 0.6 | PASS
| Col Middle (mm) | 79.63 | 79.97 | 0.34 | 0.6 | PASS
| Col Bottom (mm) | 79.97 | 80.24 | 0.27 | 0.6 | PASS
Ghosting | Signal-to-Bkgd ratio | 0.041% | 0.03% | 0.00011 | 0.5% | PASS
Spatial Resoloution Contrast Response Manual | 1.1mm Horizontal | 53.1% | 51.05% | 2.05% | 8% | PASS
| 1.0mm Horizontal | 16.1% | 14.59% | 1.51% | 8% | PASS
| 0.9mm Horizontal | 11.3% | 10.83% | 0.47% | 8% | PASS
| 0.8mm Horizontal | 1% | 2.25% | 1.25% | 8% | PASS
| 1.1mm Vertical | 52% | 50.85% | 1.15% | 8% | PASS
| 1.0mm Vertical | 20% | 19.83% | 0.17% | 8% | PASS
| 0.9mm Vertical | 7.4% | 7.3% | 0.1% | 8% | PASS
| 0.8mm Vertical | 4.7% | 1.19% | 3.51% | 8% | PASS
Spatial Resoloution Contrast Response Automatic | 1.1mm Horizontal | 69.21% | 71.8% | 2.59% | 8% | PASS
| 1.0mm Horizontal | 19.11% | 23.51% | 4.4% | 8% | PASS
| 0.9mm Horizontal | 4.21% | 5.05% | 0.84% | 8% | PASS
| 0.8mm Horizontal | 7.31% | 9.15% | 1.84% | 8% | PASS
| 1.1mm Vertical | 65.72% | 62.36% | 3.36% | 8% | PASS
| 1.0mm Vertical | 18.14% | 22.38% | 4.24% | 8% | PASS
| 0.9mm Vertical | 14.2% | 6.44% | 7.76% | 8% | PASS
| 0.8mm Vertical | 11.38% | 11.51% | 0.13% | 8% | PASS

## Raig 2 Ax T1 SE
Test Case ID | Raig2_Ax_T1_SE |  | Test Description | Initial Overall Test
Git Commit | b0a868e
Tested By | J Tracey
Test Date | 2024-09-12
|  |  | Scanner | Signa Artist
|  |  | Scan Date | 2024-05-28
|  |  | Series Description | Ax T1 SE
Test | Comparable Measurement | Software Result | Manual Result | Difference | Tolerance | PASS/ FAIL | Comments
Uniformity | Integral Uniformity | 95.51% | 96.18% | 0.6966% | 2% | PASS
Slice Position | Distance Slice 1(mm) | 1.86 | 2.06 | 0.2 | 0.6 | PASS
| Distance Slice 11(mm) | 1.27 | 1.41 | 0.14 | 0.6 | PASS
Slice Width | Distance (mm) | 5.53 | 5.95 | 0.42 | 0.6 | PASS
SNR | SNR | 324.19 | 332 | 2.352% | 3% | PASS
Geometric Accuracy MagNet Method | Row Top (mm) | 79.87 | 79.9 | 0.03 | 0.6 | PASS
| Row Middle (mm) | 79.98 | 79.9 | 0.08 | 0.6 | PASS
| Row Bottom (mm) | 79.7 | 79.7 | 0 | 0.6 | PASS
| Col Top (mm) | 79.97 | 80.4 | 0.43 | 0.6 | PASS
| Col Middle (mm) | 80.1 | 80.2 | 0.1 | 0.6 | PASS
| Col Bottom (mm) | 80.2 | 80.1 | 0.1 | 0.6 | PASS
Ghosting | Signal-to-Bkgd ratio | 0.014% | 0.03% | 0.00016 | 0.5% | PASS
Spatial Resoloution Contrast Response Manual | 1.1mm Horizontal | 95.3% | 94.2% | 1.1% | 8% | PASS
| 1.0mm Horizontal | 84.9% | 84.8% | 0.1% | 8% | PASS
| 0.9mm Horizontal | 52% | 49.67% | 2.33% | 8% | PASS
| 0.8mm Horizontal | 9.37% | 10.05% | 0.68% | 8% | PASS
| 1.1mm Vertical | 89.7% | 89.15% | 0.55% | 8% | PASS
| 1.0mm Vertical | 85.8% | 82% | 3.8% | 8% | PASS
| 0.9mm Vertical | 43.8% | 44.83% | 1.03% | 8% | PASS
| 0.8mm Vertical | 17.9% | 20.06% | 2.16% | 8% | PASS
Spatial Resoloution Contrast Response Automatic | 1.1mm Horizontal | 98.61% | 98.88% | 0.27% | 8% | PASS
| 1.0mm Horizontal | 93.49% | 87.4% | 6.09% | 8% | PASS
| 0.9mm Horizontal | 66.93% | 68.67% | 1.74% | 8% | PASS
| 0.8mm Horizontal | 18.67% | 22.07% | 3.4% | 8% | PASS
| 1.1mm Vertical | 97.65% | 94.3% | 3.35% | 8% | PASS
| 1.0mm Vertical | 95.66% | 95.24% | 0.42% | 8% | PASS
| 0.9mm Vertical | 49.8% | 51.26% | 1.46% | 8% | PASS
| 0.8mm Vertical | 20.42% | 14.48% | 5.94% | 8% | PASS

## Raig 2 Sag T1 SE
Test Case ID | Raig2_Sag_T1_SE |  | Test Description | Initial Overall Test
Git Commit | b0a868e
Tested By | J Tracey
Test Date | 2024-09-12
|  |  | Scanner | Signa Artist
|  |  | Scan Date | 2024-05-28
|  |  | Series Description | Sag T1 SE
Test | Comparable Measurement | Software Result | Manual Result | Difference | Tolerance | PASS/ FAIL | Comments
Uniformity | Integral Uniformity | 95.56% | 97.13% | 1.616% | 2% | PASS
Slice Position | Distance Slice 1(mm) | -1.07 | -0.8 | 0.27 | 0.6 | PASS
| Distance Slice 11(mm) | 1.86 | 1.63 | 0.23 | 0.6 | PASS
Slice Width | Distance (mm) | 4.96 | 5.07 | 0.11 | 0.6 | PASS
SNR | SNR | 297.9 | 305.98 | 2.641% | 3% | PASS
Geometric Accuracy MagNet Method | Row Top (mm) | 79.95 | 80.51 | 0.56 | 0.6 | PASS
| Row Middle (mm) | 79.97 | 80.35 | 0.38 | 0.6 | PASS
| Row Bottom (mm) | 79.89 | 80.18 | 0.29 | 0.6 | PASS
| Col Top (mm) | 80.12 | 79.87 | 0.25 | 0.6 | PASS
| Col Middle (mm) | 80.2 | 80.19 | 0.01 | 0.6 | PASS
| Col Bottom (mm) | 80.2 | 79.87 | 0.33 | 0.6 | PASS
Ghosting | Signal-to-Bkgd ratio | 0.037% | 0.018% | 0.00019 | 0.5% | PASS
Spatial Resoloution Contrast Response Manual | 1.1mm Horizontal | 90.2% | 91.59% | 1.39% | 8% | PASS
| 1.0mm Horizontal | 82.7% | 83.39% | 0.69% | 8% | PASS
| 0.9mm Horizontal | 50.4% | 48.88% | 1.52% | 8% | PASS
| 0.8mm Horizontal | 4.4% | 8.22% | 3.82% | 8% | PASS
| 1.1mm Vertical | 95.4% | 96.41% | 1.01% | 8% | PASS
| 1.0mm Vertical | 90.8% | 91.75% | 0.95% | 8% | PASS
| 0.9mm Vertical | 48.5% | 50.05% | 1.55% | 8% | PASS
| 0.8mm Vertical | 3.7% | 7.46% | 3.76% | 8% | PASS
Spatial Resoloution Contrast Response Automatic | 1.1mm Horizontal | 96.35% | 92.01% | 4.34% | 8% | PASS
| 1.0mm Horizontal | 97.55% | 98.1% | 0.55% | 8% | PASS
| 0.9mm Horizontal | 68.79% | 65.32% | 3.47% | 8% | PASS
| 0.8mm Horizontal | 17.95% | 19.38% | 1.43% | 8% | PASS
| 1.1mm Vertical | 98.98% | 96.79% | 2.19% | 8% | PASS
| 1.0mm Vertical | 97.8% | 96.79% | 1.01% | 8% | PASS
| 0.9mm Vertical | 56.22% | 54.78% | 1.44% | 8% | PASS
| 0.8mm Vertical | 16.59% | 17.18% | 0.59% | 8% | PASS

## Raig 1 ACR AxT2
Test Case ID | Raig_1_Ax_T2 |  | Test Description | Initial Overall Test
Git Commit | b0a868e
Tested By | J Tracey
Test Date | 2024-09-17
|  |  | Scanner | Signa Artist
|  |  | Scan Date | 2024-05-28
|  |  | Series Description | ACR AxT2
Test | Comparable Measurement | Software Result | Manual Result | Difference | Tolerance | PASS/ FAIL | Comments
Uniformity | Integral Uniformity | 76.49% | 77.04% | 0.7139% | 2% | PASS
Slice Position | Distance Slice 1(mm) | -1.76 | -1.4 | 0.36 | 0.6 | PASS
| Distance Slice 11(mm) | -1.76 | -1.41 | 0.35 | 0.6 | PASS
Slice Width | Distance (mm) | 5.15 | 5.52 | 0.37 | 0.6 | PASS
SNR | SNR | 237.48 | 237 | 0.2025% | 3% | PASS
Geometric Accuracy MagNet Method | Row Top (mm) | 79.67 | 79.6 | 0.07 | 0.6 | PASS
| Row Middle (mm) | 79.98 | 79.72 | 0.26 | 0.6 | PASS
| Row Bottom (mm) | 79.83 | 79.62 | 0.21 | 0.6 | PASS
| Col Top (mm) | 79.64 | 79.49 | 0.15 | 0.6 | PASS
| Col Middle (mm) | 79.98 | 79.89 | 0.09 | 0.6 | PASS
| Col Bottom (mm) | 80.11 | 80.04 | 0.07 | 0.6 | PASS
Ghosting | Signal-to-Bkgd ratio | 0.273% | 0.134% | 0.00139 | 0.5% | PASS
Spatial Resoloution Contrast Response Manual | 1.1mm Horizontal | 51.8% | 50.56% | 1.24% | 8% | PASS
| 1.0mm Horizontal | 58% | 56.59% | 1.41% | 8% | PASS
| 0.9mm Horizontal | 3% | 1.74% | 1.26% | 8% | PASS
| 0.8mm Horizontal | 6.9% | 6.15% | 0.75% | 8% | PASS
| 1.1mm Vertical | 19.5% | 16.9% | 2.6% | 8% | PASS
| 1.0mm Vertical | 12.5% | 12.92% | 0.42% | 8% | PASS
| 0.9mm Vertical | 9.4% | 8.24% | 1.16% | 8% | PASS
| 0.8mm Vertical | 6% | 5.93% | 0.07% | 8% | PASS
Spatial Resoloution Contrast Response Automatic | 1.1mm Horizontal | 56.71% | 55.62% | 1.09% | 8% | PASS
| 1.0mm Horizontal | 62.16% | 55.85% | 6.31% | 8% | PASS
| 0.9mm Horizontal | 19.66% | 16.09% | 3.57% | 8% | PASS
| 0.8mm Horizontal | 1.12% | 1.05% | 0.07% | 8% | PASS
| 1.1mm Vertical | 30.01% | 29.08% | 0.93% | 8% | PASS
| 1.0mm Vertical | 15.21% | 22.84% | 7.63% | 8% | PASS
| 0.9mm Vertical | 8.39% | 11.6% | 3.21% | 8% | PASS
| 0.8mm Vertical | 6.27% | 6.51% | 0.24% | 8% | PASS

## Raig 2 Cor T1 SE
Test Case ID | Raig_2_Cor_T1 |  | Test Description | Initial Overall Test
Git Commit | b0a868e
Tested By | J Tracey
Test Date | 2024-09-17
|  |  | Scanner | Signa Artist
|  |  | Scan Date | 2024-05-28
|  |  | Series Description | Cor T1 SE
Test | Comparable Measurement | Software Result | Manual Result | Difference | Tolerance | PASS/ FAIL | Comments
Uniformity | Integral Uniformity | 95.28% | 95.82% | 0.5636% | 2% | PASS
Slice Position | Distance Slice 1(mm) | -1.86 | -1.47 | 0.39 | 0.6 | PASS
| Distance Slice 11(mm) | 0.59 | 0.35 | 0.24 | 0.6 | PASS
Slice Width | Distance (mm) | 4.85 | 4.89 | 0.04 | 0.6 | PASS
SNR | SNR | 274.59 | 274.3037 | 0.1044% | 3% | PASS
Geometric Accuracy MagNet Method | Row Top (mm) | 79.83 | 79.76 | 0.07 | 0.6 | PASS
| Row Middle (mm) | 79.9 | 80.13 | 0.23 | 0.6 | PASS
| Row Bottom (mm) | 79.83 | 80.32 | 0.49 | 0.6 | PASS
| Col Top (mm) | 79.77 | 79.95 | 0.18 | 0.6 | PASS
| Col Middle (mm) | 80.27 | 80.31 | 0.04 | 0.6 | PASS
| Col Bottom (mm) | 80.17 | 80.49 | 0.32 | 0.6 | PASS
Ghosting | Signal-to-Bkgd ratio | 0.025% | 0.031% | 6e-05 | 0.5% | PASS
Spatial Resoloution Contrast Response Manual | 1.1mm Horizontal | 93.3% | 93.4% | 0.1% | 8% | PASS
| 1.0mm Horizontal | 86.6% | 87.78% | 1.18% | 8% | PASS
| 0.9mm Horizontal | 57.8% | 58.61% | 0.81% | 8% | PASS
| 0.8mm Horizontal | 3% | 5.71% | 2.71% | 8% | PASS
| 1.1mm Vertical | 92% | 92.88% | 0.88% | 8% | PASS
| 1.0mm Vertical | 87.9% | 91.31% | 3.41% | 8% | PASS
| 0.9mm Vertical | 37.2% | 33.35% | 3.85% | 8% | PASS
| 0.8mm Vertical | 10.5% | 9.83% | 0.67% | 8% | PASS
Spatial Resoloution Contrast Response Automatic | 1.1mm Horizontal | 98.14% | 98.13% | 0.01% | 8% | PASS
| 1.0mm Horizontal | 94.37% | 96.62% | 2.25% | 8% | PASS
| 0.9mm Horizontal | 64.14% | 62.45% | 1.69% | 8% | PASS
| 0.8mm Horizontal | 16.29% | 18.95% | 2.66% | 8% | PASS
| 1.1mm Vertical | 98.64% | 95.88% | 2.76% | 8% | PASS
| 1.0mm Vertical | 97.05% | 95.51% | 1.54% | 8% | PASS
| 0.9mm Vertical | 51.38% | 55.19% | 3.81% | 8% | PASS
| 0.8mm Vertical | 23.08% | 20.05% | 3.03% | 8% | PASS

## Template
Test Case ID |  |  | Test Description
Git Commit
Tested By
Test Date
Scanner
Scan Date
Series Description
Test | Comparable Measurement | Software Result | Manual Result | Difference | Tolerance | PASS/ FAIL | Comments
Uniformity | Integral Uniformity |  |  | Not Run | 2% | Not Run
Slice Position | Distance Slice 1(mm) |  |  | Not Run | 0.6 | Not Run
| Distance Slice 11(mm) |  |  | Not Run | 0.6 | Not Run
Slice Width | Distance (mm) |  |  | Not Run | 0.6 | Not Run
SNR | SNR |  |  | Not Run | 8% | Not Run
Geometric Accuracy MagNet Method | Row Top (mm) |  |  | Not Run | 0.6 | Not Run
| Row Middle (mm) |  |  | Not Run | 0.6 | Not Run
| Row Bottom (mm) |  |  | Not Run | 0.6 | Not Run
| Col Top (mm) |  |  | Not Run | 0.6 | Not Run
| Col Middle (mm) |  |  | Not Run | 0.6 | Not Run
| Col Bottom (mm) |  |  | Not Run | 0.6 | Not Run
Ghosting | Signal-to-Bkgd ratio |  |  | Not Run | 0.5% | Not Run
Spatial Resoloution Contrast Response Manual | 1.1mm Horizontal |  |  | Not Run | 8% | Not Run
| 1.0mm Horizontal |  |  | Not Run | 8% | Not Run
| 0.9mm Horizontal |  |  | Not Run | 8% | Not Run
| 0.8mm Horizontal |  |  | Not Run | 8% | Not Run
| 1.1mm Vertical |  |  | Not Run | 8% | Not Run
| 1.0mm Vertical |  |  | Not Run | 8% | Not Run
| 0.9mm Vertical |  |  | Not Run | 8% | Not Run
| 0.8mm Vertical |  |  | Not Run | 8% | Not Run
Spatial Resoloution Contrast Response Automatic | 1.1mm Horizontal |  |  | Not Run | 8% | Not Run
| 1.0mm Horizontal |  |  | Not Run | 8% | Not Run
| 0.9mm Horizontal |  |  | Not Run | 8% | Not Run
| 0.8mm Horizontal |  |  | Not Run | 8% | Not Run
| 1.1mm Vertical |  |  | Not Run | 8% | Not Run
| 1.0mm Vertical |  |  | Not Run | 8% | Not Run
| 0.9mm Vertical |  |  | Not Run | 8% | Not Run
| 0.8mm Vertical |  |  | Not Run | 8% | Not Run